## Slide 1: Advancing PEPFAR Program Quality Assurance Through HIV Rapid Diagnostic Test Lot Testing
USAID Global Health Supply Chain-Quality Assurance Program
August 2024
Ashini Fernando, PhD, MPH (FHI 360)
David Jenkins, PhD (FHI 360)
Roger Peck (PATH)

## Slide 2: Agenda
Introduction to USAID Global Health Supply Chain-Quality Assurance Program (GHSC-QA)
GHSC HIV RDT lot testing program
HIV RDT lot testing standardization
HIV RDT testing outcomes

## Slide 3: USAID Global Health Supply Chain-Quality Assurance Program (GHSC-QA)
FHI 360’s Product Quality and Compliance (PQC) department provides independent and objective quality assurance and quality control services to clients.
PQC has staff in the United States, Thailand, and South Africa; ISO 17025-accredited testing laboratories in United States and Thailand; ISO 9001-certified quality management system.
Through GHSC-QA, FHI 360 delivers comprehensive quality assurance services employing a risk-based approach to ensure the safety, efficacy, and quality of health commodities procured on behalf of USAID.
GHSC-QA contract objectives include:
  Quality assurance
  Quality control
  Technical assistance
  Global collaborations
USAID Global Health Supply Chain Program

## Slide 4: GHSC HIV RDT Lot Testing Program
Procurement groups
Notification to FHI 360 of
Sampling agents
Notification of HIV RDT orders (and incidents)
3 testing laboratories
Result reporting
within 6 weeks
Coordinate with
sampling agents
Test method development
Sampling strategy
Product manufacturing consistency monitoring
Incident management
Communicate results and next steps
to procurement groups
200 HIV RDT samples mailed
WHO-prequalified
HIV RDTS
1
2
3
4
5
GHSC-QA

## Slide 5: HIV RDT Testing Standardization
Guided by ISO 2859-1:1999 Sampling Procedures for Inspection by Attributes
Key questions addressed:
  How many HIV RDTs to test? Sample size
  What are the attributes applied to accept or reject a lot? Nonconformant results
  What should be the percent level of defects in a lot that is tolerated by a user? Acceptance quality limits (AQL)
      Factors considered: stringency, WHO guidance on HIV RDT sensitivity and specificity, waste
  How should we qualify specimen panels?
      Factors considered: specimen types, availability, cost, and storage
  How should the RDTs be monitored? Sampling plan
Notes: :

## Slide 6
Sample Size and Acceptance Quality Limit (AQL) Determination Based on ISO 2859-1
Stringency
Stringency

## Slide 7: Optimal AQL and Testing Level for HIV RDTs: 0.65, S-4
Sampling plan
80 HIV-positive unique specimens with 80 HIV RDTs
80 HIV-negative unique specimens with 80 HIV RDTs
AQL: accept 1 nonconformant
reject 2 nonconformants
Alternative sampling plan
20 positive specimens in quadruplicate with 80 RDTs
20 negative specimens in quadruplicate with 80 RDTs
AQL: accept 1 nonconformant
reject 2 nonconformants
PloS One, 18(4), e0284175. https://doi.org/10.1371/journal.pone.0284175

## Slide 8: A Snapshot of HIV Specimen Panel Qualification Outcomes With 5 HIV RDT Brands
80 HIV-positive specimens and 80 HIV-negative specimens
Each specimen in 4 replicate/panel
3 lots per brand of HIV RDTs tested
Reactive
Non-reactive

## Slide 9: HIV RDT Sampling Plan
*Qualification score determination
For 80 positives and 80 negatives; AQL 0.65, 1 accept/2 reject
  If a lot is accepted with 0 nonconformities, add 5 to the qualification score
  If a lot is accepted with 1 nonconformity, add 1 to the qualification score
  Any lot not accepted results in the qualification score being reset to 0
 | Qualification score* | Minimum lots tested to reach next levels | Test frequency | Sample size
Testing level 1 | - | 15 | 100% | 80 positive and 80 negative
Testing level 2 | 75 | 50 | 100% | 80 positive and 80 negative
Testing level 3 | 250 | 75 | Skip lot 1 in 4 | 80 positive and 80 negative
Testing level 4 | 375 | - | 1 per quarter | 20 positive and 20 negative
ISO 2859-3: Skip lot sampling procedure
USAID Global Health Supply Chain Program

## Slide 10: HIV and HIV/Syphilis RDT Lot Testing Outcomes
Chart: HIV RDTs, 22 brands
| Lots tested | Lots not tested
 | 1146 | 1356
Lots tested
Lots not
tested
7 Lots
rejected
I
incident
Total lots procured 2,502
Chart: HIV/Syphilis combo, 2 brands
| Lots tested | Lots not tested
 | 62 | 23
Lots not
tested
Total lots procured 85

## Slide 11: Investigation Resulted in Improvements to the Instructions for Use for a HIV Rapid Diagnostic Test
Product was quarantined after yielding false positive results.
Testing through GHSC-QA and by the manufacturer found a higher rate of invalid results occurred when the RDT was used directly on a laboratory bench or a plastic surface, but invalid results did not occur when tests were used on a paper-backed surface. This phenomenon is attributed to static electricity that affects the volume of buffer dispensed in a droplet from the buffer bottle.
As a result of the investigation:
  The manufacturer prepared updates to the instructions for use to read that testing be performed on a paper-based surface.
  The manufacturer submitted a notification of labeling change to WHO Prequalification. This labeling change will safeguard future procurements that use this HIV RDT as a component of their testing algorithm.
FHI 360 and PATH worked closely and communicated regularly with global stakeholders including WHO, USAID, the Global Fund, the USAID Global Health Supply Chain-Technical Assistance (GHSC-TA) program, and the USAID Global Health Supply Chain-Rapid Test Kit (GHSC-RTK) program. FHI 360 also maintained a positive collaboration with the manufacturer as each entity worked to investigate the concern of invalid results.

## Slide 12: A Smaller Scale Lot Testing Method
Use cases: verification of RDT performance prior to use (new brands, new lots), routine QC checks, temperature excursions, IVD testing at sentinel sites, troubleshooting, EQA
Potential sample sizes and AQLs
  20 sample size with an AQL of 0.65 (accept/reject ratio 0/1)
  13 samples size with an AQL of 1.0 (accept/reject ratio 0/1)
Potential QC material
  Strong positive, medium positive(s), weak positive(s)
  Negative(s), a negative + specificity substance 1, a negative + specificity substance 2
USAID Global Health Supply Chain Program

## Slide 13: Acknowledgments
This work was supported through the United States Agency for International Development Contract No. AID-OAA-C-15-00001
FHI 360
  David Jenkins, PhD
  Mary Beth Sowers
  Michael Payne
  Arno Hensens
  Jeffery Tremelling
PATH
  Roger Peck
  Becky Barney
University of Maryland, Baltimore
  Niel Constantine, PhD
World Health Organization
  Anita Sands

## Slide 14: Thank you!
Contact information
Ashini Fernando: afernando@fhi360.org
David Jenkins: djenkins@fhi360.org